Title: Understanding and Managing Flow Errors on the Fresenius 2008T with Bibag
Prepared by: Kayla Askins – Biomedical Equipment Technician I
Fresenius Medical Care – Fort Lauderdale Region

### PowerPoint Deck (Training Presentation)

Slide 1 – Title Slide
Understanding and Managing Low Flow & High Flow Errors on the Fresenius 2008T with Bibag
Prepared by: Kayla Askins – Biomedical Equipment Technician I

Slide 2 – Why This Matters

- Stable flow ensures patient safety and dialysis efficiency.

- Flow errors impact conductivity, temperature, and treatment adequacy.

- Many can be resolved at the chairside by clinical staff.

Slide 3 – What the Machine Is Telling You

| Alarm | What It Means | What It Affects |
| --- | --- | --- |
| Low Flow | Dialysate not circulating fast enough | Conductivity & Temperature |
| High Flow | Excess flow or pressure difference | TMP & UF control |

Slide 4 – Common Causes on the Floor

- Air in dialysate loop (Bibag not seated, incomplete rinse).

- Kinked or pinched acid/bicarb lines.

- Bibag vent filter clogged or wet.

- Weak RO water pressure or closed loop valve.

- UF pump imbalance.

Slide 5 – What to Do During Treatment

- Stay with the patient; silence the alarm.

- Check behind machine for kinks or pinches.

- Inspect Bibag—should deflate smoothly.

- Observe RO flow (other machines alarming?).

- If alarm repeats → pause treatment and notify Biomed.

Slide 6 – Real Clinical Scenarios

- Multiple chairs alarming → RO pressure drop.

- Single machine alarming → Bibag vent blocked.

- Flow unstable → UF pump overpulling.

Slide 7 – Prevention Steps

| Task | Performed By | Purpose |
| --- | --- | --- |
| Full rinse before treatment | Clinical | Removes air |
| Replace Bibag filter daily | Clinical | Prevents vacuum lock |
| Verify RO pressure | Clinical & Biomed | Ensures steady supply |
| Calibrate deaeration pump | Biomed | Maintains true flow |
| Flush heater quarterly | Biomed | Prevents scale buildup |

Slide 8 – Key Takeaway

"When the 2008T alarms, it’s protecting the patient — not failing the treatment."
Collaboration between nursing and Biomed prevents downtime and ensures patient safety.

Slide 9 – Contact & QR
Scan the QR Code for full troubleshooting guide and updates.
Prepared by: Kayla Askins – Biomedical Equipment Technician I
Fresenius Medical Care – Fort Lauderdale Region

### Laminated Quick Reference (PDF)

Front (Clinical Side)
Low & High Flow Error Quick Check
1️⃣ Verify acid & bicarb lines – no kinks.
2️⃣ Confirm Bibag seated, not inflated.
3️⃣ Observe Flow (FLWP ≈ 800 mL/min).
4️⃣ If multiple alarms → contact Biomed.

Common Floor Causes

- Air in dialysate circuit.

- Low RO pressure.

- Bibag vent filter blocked.

- Twisted or disconnected concentrate lines.

Diagram: Color-coded flow map
(Blue = Dialysate Flow | Orange = UF | Yellow = RO Feed)

QR Code: Scan for full troubleshooting guide.

Back (Biomed Reference)

- Verify deaeration pump calibration.

- Check UF synchronization.

- Inspect valves V31–V38.

- Flush heater/heat exchanger.

- Review calibration logs.

Contact Information
Kayla Askins – Biomedical Equipment Technician I
Fresenius Medical Care – Fort Lauderdale Region
📧 kayla.askins@fmcna.com | ☎️ (xxx) xxx-xxxx

Design: Fresenius blue/gray theme with QR code and branding footer.

Title: Understanding and Managing Flow Errors on the Fresenius 2008T with Bibag
Prepared by: Kayla Askins – Biomedical Equipment Technician I
Fresenius Medical Care – Fort Lauderdale Region

### PowerPoint Deck (Training Presentation)

Slide 1 – Title Slide
Understanding and Managing Low Flow & High Flow Errors on the Fresenius 2008T with Bibag
Prepared by: Kayla Askins – Biomedical Equipment Technician I

Slide 2 – Why This Matters

- Stable flow ensures patient safety and dialysis efficiency.

- Flow errors impact conductivity, temperature, and treatment adequacy.

- Many can be resolved at the chairside by clinical staff.

Slide 3 – What the Machine Is Telling You

| Alarm | What It Means | What It Affects |
| --- | --- | --- |
| Low Flow | Dialysate not circulating fast enough | Conductivity & Temperature |
| High Flow | Excess flow or pressure difference | TMP & UF control |

Slide 4 – Common Causes on the Floor

- Air in dialysate loop (Bibag not seated, incomplete rinse).

- Kinked or pinched acid/bicarb lines.

- Bibag vent filter clogged or wet.

- Weak RO water pressure or closed loop valve.

- UF pump imbalance.

Slide 5 – What to Do During Treatment

- Stay with the patient; silence the alarm.

- Check behind machine for kinks or pinches.

- Inspect Bibag—should deflate smoothly.

- Observe RO flow (other machines alarming?).

- If alarm repeats → pause treatment and notify Biomed.

Slide 6 – Real Clinical Scenarios

- Multiple chairs alarming → RO pressure drop.

- Single machine alarming → Bibag vent blocked.

- Flow unstable → UF pump overpulling.

Slide 7 – Prevention Steps

| Task | Performed By | Purpose |
| --- | --- | --- |
| Full rinse before treatment | Clinical | Removes air |
| Replace Bibag filter daily | Clinical | Prevents vacuum lock |
| Verify RO pressure | Clinical & Biomed | Ensures steady supply |
| Calibrate deaeration pump | Biomed | Maintains true flow |
| Flush heater quarterly | Biomed | Prevents scale buildup |

Slide 8 – Key Takeaway

"When the 2008T alarms, it’s protecting the patient — not failing the treatment."
Collaboration between nursing and Biomed prevents downtime and ensures patient safety.

Slide 9 – Contact & QR
Scan the QR Code for full troubleshooting guide and updates.
Prepared by: Kayla Askins – Biomedical Equipment Technician I
Fresenius Medical Care – Fort Lauderdale Region

### Laminated Quick Reference (PDF)

Front (Clinical Side)
Low & High Flow Error Quick Check
1️⃣ Verify acid & bicarb lines – no kinks.
2️⃣ Confirm Bibag seated, not inflated.
3️⃣ Observe Flow (FLWP ≈ 800 mL/min).
4️⃣ If multiple alarms → contact Biomed.

Common Floor Causes

- Air in dialysate circuit.

- Low RO pressure.

- Bibag vent filter blocked.

- Twisted or disconnected concentrate lines.

Diagram: Color-coded flow map
(Blue = Dialysate Flow | Orange = UF | Yellow = RO Feed)

QR Code: Scan for full troubleshooting guide.

Back (Biomed Reference)

- Verify deaeration pump calibration.

- Check UF synchronization.

- Inspect valves V31–V38.

- Flush heater/heat exchanger.

- Review calibration logs.

Contact Information
Kayla Askins – Biomedical Equipment Technician I
Fresenius Medical Care – Fort Lauderdale Region
📧 kayla.askins@fmcna.com | ☎️ (xxx) xxx-xxxx

Design: Fresenius blue/gray theme with QR code and branding footer.