Excellent, Kayla — here’s a fully structured and professional clinical email written in your voice, reflecting the actual field findings you’ve been uncovering in your diagnostic work on the Fresenius 2008T.

This message is ready to send to your clinical team and includes:

- Clear technical context for staff

- Probable causes discovered from your investigation

- Step-by-step clinical troubleshooting guide (with referenced figures)

- Prevention and escalation instructions

### To: Clinical Team – Imperial Pointe & Hampton Clinics

### From: Kayla Askins, Biomedical Equipment Technician I

### Subject: Resolving Flow and Conductivity Issues on Fresenius 2008T Machines with Bibag

### Opening Paragraph

Team,

The purpose of this email is to address the flow and conductivity issues we experienced today on several Fresenius 2008T dialysis machines equipped with Bibag systems. These alarms occurred intermittently during treatment, presenting as “Low Flow,” “Flow Error,” and “Conductivity Self-Test Fail” messages.

After reviewing the affected stations and analyzing debug values, the probable cause appears to be air entrapment and restricted dialysate flow within the hydraulic circuit — specifically linked to Bibag venting issues, deaeration pump performance, or low water supply pressure from the RO feed. These factors prevent the machine from maintaining steady dialysate flow and accurate conductivity readings.

### Body Paragraph 1 – Plan of Action: Troubleshooting Steps

To minimize treatment interruptions and assist in identifying the root cause quickly, please follow the steps below when this occurs.

#### Step 1 – Check the Acid and Bicarb Lines (Figure 1)

First, visually inspect all acid and bicarbonate lines behind the machine.

- Ensure that each line is fully connected, untwisted, and not kinked.

- If you notice restricted movement or tension, gently reposition the line.

🖼️ Figure 1: Hose Routing Behind 2008T Machine
(Highlighted arrows indicate proper routing for acid (red) and bicarb (blue) lines. Boxes mark common pinch points.)

#### Step 2 – Inspect the Bibag (Figure 2)

Next, check the Bibag on the front panel:

- Confirm the Bibag is properly seated on the connector port.

- Observe whether it inflates or remains puffed up after use.

  - A Bibag that fails to deflate usually indicates a blocked vent filter or valve (V105).

🖼️ Figure 2: Proper Bibag Seating and Deflation
(Green check = correct placement; red X = overinflated bag indicating vent restriction.)

If this occurs:

- Replace the Bibag vent filter.

- Gently tap the Bibag port to ensure the valve is not stuck closed.

#### Step 3 – Confirm Water Supply and Rinse (Figure 3)

Then, verify that RO water pressure is stable:

- Look for a steady “RO Pressure OK” indicator on the screen.

- Start a Rinse cycle and observe the Flow (FLWP) reading on the TMP/Low Soft screen.

  - Normal range: ~800 mL/min

  - Below 700 mL/min suggests low water pressure or air in the loop.

🖼️ Figure 3: TMP/Low Soft Screen
(Highlighted box around “FLWP” displays actual flow rate reading. Arrows show stable vs unstable flow patterns.)

If flow is low during rinse:

- Allow rinse to continue for 3–5 minutes.

- If it stabilizes, return to treatment.

- If not, pause and call Biomed.

#### Step 4 – Observe for Conductivity Stability (Figure 4)

After rinsing, monitor the conductivity value on the main screen:

- It should rise steadily between 13.0 – 14.0 mS/cm.

- Fluctuations indicate an ongoing mixing or flow issue.

🖼️ Figure 4: Conductivity Display Reading During Rinse
(Arrows show steady vs unstable conductivity trend line.)

If conductivity does not stabilize:

- Replace Bibag and inspect for wet vent filters or air in the deaeration pump.

- Do not proceed with treatment until Biomed evaluates the unit.

### Body Paragraph 2 – Preventive Measures and Troubleshooting Tips

To prevent recurrence and ensure smoother treatment setups:

- Run full Rinse cycles before connecting patients.

- Replace Bibag filters daily, especially if they appear wet or powder-coated.

- Avoid twisting or stretching acid and bicarb lines when rolling machines.

- Verify RO pressure each morning at the main RO console (18–22 psi normal).

- If multiple machines alarm at once, pause treatments and notify Biomed immediately — this suggests an RO or loop-wide issue, not individual machine failure.

Troubleshooting Summary Table

| Symptom | Likely Cause | Quick Action |
| --- | --- | --- |
| Low Flow Error | Air in loop or restricted line | Check hoses, rinse, verify RO |
| Flow Error During Treatment | Bibag vent blocked or wet filter | Replace filter, ensure Bibag deflates |
| Conductivity Self-Test Fail | Air or unstable temperature | Allow rinse cycle; confirm flow stabilizes |
| Repeated Alarms on Multiple Units | RO feed pressure low | Notify Biomed immediately |

### Closing Paragraph

In summary, the flow and conductivity issues observed today are likely the result of air entrapment, Bibag venting restrictions, or low RO pressure. By performing the above visual checks and rinse verification steps, most alarms can be resolved at the chairside without interrupting treatment.

If any machine continues to alarm after these steps:

- Pause treatment to ensure patient safety.

- Record the station number, time, and what was observed (e.g., Bibag not deflating, low flow reading).

- Contact Biomed immediately for further diagnostics.

Thank you all for your teamwork and attention to these details. Your proactive checks make a major difference in maintaining safe and efficient dialysis treatments for our patients.

Best regards,
Kayla Askins
Biomedical Equipment Technician I
Fresenius Medical Care – Fort Lauderdale Region
📧 kayla.askins@fmcna.com | ☎️ (xxx) xxx-xxxx

Would you like me to add the labeled figure set (Figures 1–4) as visual attachments to this email — with the arrows, color highlights, and short captions so it’s presentation-ready for your next shift huddle?